## CAN Node ID
Priority1 | Node ID (lower 5 bits of Message ID) decimal | Software Part Number | Description
Highest Priority | 00
01
02
03
04
05
06
| 07 | Sterilization Control
08
| 09 | Display
| 10 | Test Interface
11
12
13
14
15
16
| 17 | Connectivity Interface
18
19
20
21
22
23
24
25
26
27
28
29
| 30 | Sensor Control
31
Lowest Priority | 254 | Firmware Update Tool
1) Priority is based upon CAN hardware when the Function Code and Node ID are combined to form the CAN Message ID

## CAN Frames
FC 0: Message Request
Extended ID: | 0x1000XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x000
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 00
Frame: | 1 of 1
Description: | Message Request - This is to be utilized to request information from another node, or all nodes if the Node ID is 0xFF.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Node ID - the node(s) info is to come from (0xFF represents all Nodes)
Data 2 | Function Code - the data being requested (Bits 0-7)
Data 3 | Function Code - the data being requested (Bits 8-15)
FC 1: Message Reception Status
Extended ID: | 0x1001XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x001
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x4
Function Code: | 01
Frame: | 1 of 1
Description: | Message Reception Status - When a message is received into a node, a status of the message is to be returned to the node that sent the information. Refer to Message Receptions tab for the Nodes and IDs that are to have this message sent.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Node ID - the node the info was from
Data 2 | Function Code - the function code the node received (Bits 0-7)
Data 3 | Function Code - the function code the node received (Bits 8-15)
Data 4 | Status of Reception (ACK=0, NACK=1)
FC 2: Active/Recently Inactive Error
Extended ID: | 0x1002XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x002
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC | 0x8
Function Code: | 02
Frame: | 1 of 1
Description: | Error Notification - An error code with time/date stamp that has had a status change.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code (Bits 0-7)
Data 2 | Error Code (Bits 8-15)
Data 3 | Error Code (Bits 16-23)
Data 4 | Error Code (Bits 24-31)
Data 5 | Time/Date (Bit 0-7)
Data 6 | Time/Date (Bits 8-15)
Data 7 | Time/Date (Bits 16-23)
Data 8 | Time/Date (Bits 24-31)
FC 3: Active Errors
Extended ID: | 0x1003XY85
Urgency: | 0x1 (Normal)
Function Code: | 0x003
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Function Code: | 3
Frame: | 1 of 5
Description: | An array of active errors.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 1 (Bits 0-7)
Data 2 | Error Code 1 (Bits 8-15)
Data 3 | Error Code 1 (Bits 16-23)
Data 4 | Error Code 1 (Bits 24-31)
Data 5 | Time/Date1 (Bit 0-7)
Data 6 | Time/Date1 (Bits 8-15)
Data 7 | Time/Date1 (Bits 16-23)
Data 8 | Time/Date1 (Bits 24-31)
Extended ID: | 0x1003XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x003
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 03
Frame: | 2 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 2 (Bits 0-7)
Data 2 | Error Code 2 (Bits 8-15)
Data 3 | Error Code 2 (Bits 16-23)
Data 4 | Error Code 2 (Bits 24-31)
Data 5 | Time/Date 2 (Bit 0-7)
Data 6 | Time/Date 2 (Bits 8-15)
Data 7 | Time/Date 2 (Bits 16-23)
Data 8 | Time/Date 2 (Bits 24-31)
Extended ID: | 0x1003XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x003
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 03
Frame: | 3 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 3 (Bits 0-7)
Data 2 | Error Code 3 (Bits 8-15)
Data 3 | Error Code 3 (Bits 16-23)
Data 4 | Error Code 3 (Bits 24-31)
Data 5 | Time/Date 3 (Bit 0-7)
Data 6 | Time/Date 3 (Bits 8-15)
Data 7 | Time/Date 3 (Bits 16-23)
Data 8 | Time/Date 3 (Bits 24-31)
Extended ID: | 0x1003XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x003
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 03
Frame: | 4 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 4 (Bits 0-7)
Data 2 | Error Code 4 (Bits 8-15)
Data 3 | Error Code 4 (Bits 16-23)
Data 4 | Error Code 4 (Bits 24-31)
Data 5 | Time/Date 4 (Bit 0-7)
Data 6 | Time/Date 4 (Bits 8-15)
Data 7 | Time/Date 4 (Bits 16-23)
Data 8 | Time/Date 4 (Bits 24-31)
Extended ID: | 0x1003XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x003
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 03
Frame: | 5 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 5 (Bits 0-7)
Data 2 | Error Code 5 (Bits 8-15)
Data 3 | Error Code 5 (Bits 16-23)
Data 4 | Error Code 5 (Bits 24-31)
Data 5 | Time/Date 5 (Bit 0-7)
Data 6 | Time/Date 5 (Bits 8-15)
Data 7 | Time/Date 5 (Bits 16-23)
Data 8 | Time/Date 5 (Bits 24-31)
FC 4: Cycle Action
Extended ID: | 0x1004XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x004
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x7
Function Code: | 04
Frame: | 1 of 1
Description: | Cycle Action - Used as a command to start/stop a cycle(set), also used as a response/confirmation to the command(broadcast). If "Stop Error" is used, the Error code shall be the error that is causing the cycle to stop. Otherwise Error code shall be 0x0
Origin | Display Board, Test Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | #REF!
Data 3 | Action Flag (Broadcast=0, Set = 2)
Data 4 | Error Code (Bits 0-7)
Data 5 | Error Code (Bits 8-15)
Data 6 | Error Code (Bits 16-23)
Data 7 | Error Code (Bits 24-31)
FC 5: Cycle Parameters
Extended ID: | 0x1005XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x005
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x7
Function Code: | 05
Frame: | 1 of 1
Description: | Cycle Parameters -
Origin | Sterilization Control, Display Board, Test Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Sterilization Temperature Index (Bits 0-7) [ 0 - 25 = 250 - 275 deg F in one deg increments]
Data 3 | Sterilization Temperature Index (Bits 8 - 15)
Data 4 | Sterilization Time in Minutes
Data 5 | Default Dry Time in Minutes (Used by Wrapped, Delicate, and Unwrapped Cycle Types)
Data 6 | Dry Time in Minutes (Used by all Cycle Types)
Data 7 | #REF! | Unused | Unused | Unused | Unused | Unused | Unused | #REF!
Data 8 | #REF!
FC 6: Cycle Result
Extended ID: | 0x1006XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x006
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 06
Frame: | 1 of 1
Description: | Cycle Result- Broadcasted to indicate that a cycle has stopped and gives the reason why the cycle was stopped
Origin | Sterilization Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
FC 7: Cycle Status
Extended ID: | 0x1007XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x007
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x8
Function Code: | 07
Frame: | 1 of 1
Description: | Cycle Status - status of running cycle, sent every second during a cycle
Origin | Sterilization Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | #REF! | Unused | Unused | Unused | #REF!
Data 3 | Time Remaining in Seconds (Bits 0-7)
Data 4 | Time Remaining in Seconds (Bits 8-15)
Data 5 | Temperature (Tenths of °F) (Bits 0-7)
Data 6 | Temperature (Tenths of °F) (Bits 8-15)
Data 7 | Pressure (Tenths of PSI) (Bits 0-7)
Data 8 | Pressure (Tenths of PSI) (Bits 8-15)
FC 8: Current Date/Time
Extended ID: | 0x1008XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x008
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x5
Function Code: | 08
Frame: | 1 of 1
Description: | Current Date/Time - The current date/time kept by the Real-Time clock (RTC). This message is utilized when the date/time needs to be set on the display. This date/time is the raw seconds since 01/01/2016 @ 00:00:00.
Origin | Connectivity Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Date/Time Stamp (Bits 0-7)
Data 2 | Date/Time Stamp (Bits 8-15)
Data 3 | Date/Time Stamp (Bits 16-23)
Data 4 | Date/Time Stamp (Bits 24-31)
Data 5 | #REF!
FC 9: Switch Status
Extended ID: | 0x1009XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x009
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 09
Frame: | 1 of 1
Description: | Switch Status - The current status of the components that are a part of the sterilization control subsystem.
Origin | Sterilization Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Not Used | Not Used | Thermostat (Open=0, Closed=1) | Not Used | Door Position Switch (Dry position=0, Not Dry pos = 1) | Chamber Water Full (Empty=0, Full=1) | Door Interlock Switch (Open=0, Closed=1) | Not Used
Data 2 | Not Used | Not Used | Door Open Motor (Off=0, On=1) | Heater Relay (Off=0, On=1) | Fill Relay (Open=0, Closed=1) | Vent Relay (Open=0, Closed=1) | Air Relay (Open=0, Closed=1) | Not Used
Data 3 | Not Used | Not Used | Not Used | Config4 (Off=1, On=0) | Config3 (Off=1, On=0) | Config2 (Off=1, On=0) | Config1 (Off=1, On=0) | Config0 (Off=1, On=0)
FC 11: Cycle Count
Extended ID: | 0x100BXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x00B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x4
Function Code: | 11
Frame: | 1 of 1
Description: | Cycle Count- Provides the current total and daily cycle count, can be requested, or spontaneously when incremented.
Origin | Sterilization control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Total Cycle Count (bits 0-7)
Data 2 | Total Cycle Count (bits 8-15)
Data 3 | Daily Cycle Count (bits 0-7)
Data 4 | #REF!
FC 13: Configuration Information
Extended ID: | 0x100DXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x00D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 13
Frame: | 1 of 1
Description: | Configuration Information - The configuration of the PCB inclusive of the PCB operating mode and product type.
Origin | Sterilizer Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Operating Mode
Data 2 | Product Type
FC 14: Diagnostic Information
Extended ID: | 0x100EXY8Z (Multi-frame) OR 0x100EXY00 (Single-frame)
Urgency: | 0x1 (Normal)
Function Code: | 0x00E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | If Requesting Information = 0x0 (Single Frame); If Sending Information = 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0xZ (Number of Frames to be Sent) OR 0x0 (Single Frame Message)
DLC: | 0xY (Vaires upon specific functionality)
Function Code: | 14 (multi-frame)
Frame: | 1 of Z (no more than the maximum allowed frames by the CAN Handler)
Description: | Diagnostic Information - The diagnostic information associated with the node. The information is node dependent and is described in the individual node's Design Document. The Diagnostic Information is to be requested using the single frame version, but data for the request may come back as a single frame or a multi-frame message.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Request = 1 Data = 0 | Type of Information
Data 2 | Data (Dependent upon Node) OR Node ID if the message is a request type
Data 3 | Data (Dependent upon Node)
Data 4 | Data (Dependent upon Node)
Data 5 | Data (Dependent upon Node)
Data 6 | Data (Dependent upon Node)
Data 7 | Data (Dependent upon Node)
Data 8 | Data (Dependent upon Node)
Extended ID: | 0x100EXYC(Z-1)
Urgency: | 0x1 (Normal)
Function Code: | 0x00E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | If Requesting Information = 0x0 (Single Frame); If Sending Information = 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0xZ -1
DLC: | 0xY (Vaires upon specific functionality)
Function Code: | 14 (multi-frame)
Frame: | Z of Z (no more than the maximum allowed frames by the CAN Handler)
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Data (Dependent upon Node)
Data 2 | Data (Dependent upon Node)
Data 3 | Data (Dependent upon Node)
Data 4 | Data (Dependent upon Node)
Data 5 | Data (Dependent upon Node)
Data 6 | Data (Dependent upon Node)
Data 7 | Data (Dependent upon Node)
Data 8 | Data (Dependent upon Node)
FC 16: Product Information
Extended ID: | 0x1010XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x010
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x8
Function Code: | 16
Frame: | 1 of 1
Description: | Product Information - The product model and serial number each as 32-bit decimal numbers.
Origin | Sterilization Control, Test Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Product Serial Number (Bits 0-7)
Data 3 | Product Serial Number (Bits 8-15)
Data 4 | Product Serial Number (Bits 16-23)
Data 5 | Product Serial Number (Bits 24-31)
Data 6 | Product Model Number (EX: 050, 051 etc.)
Data 7 | Chamber Diameter in Inches (9 or 11)
Data 8 | Unused | Unused | Unused | Unused | Unused | Unused | Unused | Unused
FC 19: Product Name
Extended ID: | 0x1013XY84
Urgency: | 0x1 (Normal)
Function Code: | 0x013
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 19
Frame: | 1 of 4
Description: | Product Name - The hexadecimal ASCII characters that represent the name of the product.
Origin | Display, Connectivity Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Name Char 1
Data 2 | Product Name Char 2
Data 3 | Product Name Char 3
Data 4 | Product Name Char 4
Data 5 | Product Name Char 5
Data 6 | Product Name Char 6
Data 7 | Product Name Char 7
Data 8 | Product Name Char 8
Extended ID: | 0x1013XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x013
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 19
Frame: | 2 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Name Char 9
Data 2 | Product Name Char 10
Data 3 | Product Name Char 11
Data 4 | Product Name Char 12
Data 5 | Product Name Char 13
Data 6 | Product Name Char 14
Data 7 | Product Name Char 15
Data 8 | Product Name Char 16
Extended ID: | 0x1013XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x013
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 19
Frame: | 3 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Name Char 17
Data 2 | Product Name Char 18
Data 3 | Product Name Char 19
Data 4 | Product Name Char 20
Data 5 | Product Name Char 21
Data 6 | Product Name Char 22
Data 7 | Product Name Char 23
Data 8 | Product Name Char 24
Extended ID: | 0x1013XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x013
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x7
Function Code: | 19
Frame: | 4 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Name Char 25
Data 2 | Product Name Char 26
Data 3 | Product Name Char 27
Data 4 | Product Name Char 28
Data 5 | Product Name Char 29
Data 6 | Product Name Char 30
Data 7 | Product Name Char 31
FC 20: Clear Memory
Extended ID: | 0x1014XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x014
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 20
Frame: | 1 of 1
Description: | Clear Memory - Information is stored to memory that may need to be cleared such that defaults can be obtained. This allows for the memory of individual or all of the nodes' memory to be cleared. Once the node's memory clear operation is complete, the status is returned in this message to the node that initiated the clearing of memory.
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Node ID Message is For (0xFF represents all Nodes)
Data 2 | #REF!
Data 3 | #REF!
FC 21: Buzzer Command
Extended ID: | 0x1015XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x015
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 21
Frame: | 1 of 1
Description: | Buzzer Command- Controls buzzer on Sensor Control
Origin | Sterilization Control, Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | #REF!
FC 23: Mode Change Request
Extended ID: | 0x1017XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x017
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 23
Frame: | 1 of 1
Description: | Mode change request
Origin | Any - Node requesting the mode change
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Target Node ID (Sterilization Control=7, Display=9, Sensor Control=30, All=255)
Data 2 | Mode (Ready=1, FW Update=8, Diagnostic=12)
Data 3 | Action Flag (Broadcast=0, Set=2, Soft-Set=3)
FC 24: System Operating Mode Status
Extended ID: | 0x1018XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x018
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 24
Frame: | 1 of 1
Description: | System Operating Mode
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 |  |  |  |  | #REF!
FC 28: Reset Reason
Extended ID: | 0x101CXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x01C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 28
Frame: | 1 of 1
Description: | Reset reason- reason for the last device reset
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Power-On Reset | External Pin Reset | Watchdog | Unused | Loss-of-Lock Reset | Loss-of-Clock Reset | Low-Voltage Detect Reset | Low Leakage Wakeup Reset
Data 2 | Unused | Unused | Stop Mode Acknowledge Error Reset | EzPort Reset | MDM-AP System Request Reset | Software | Core Lockup | JTAG Generated Reset
FC 29: Product Location
Extended ID: | 0x101DXY84
Urgency: | 0x1 (Normal)
Function Code: | 0x01D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 29
Frame: | 1 of 4
Description: | Product Location- The product location as set by the user
Origin | Display, Connectivity Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Location Char 1
Data 2 | Product Location Char 2
Data 3 | Product Location Char 3
Data 4 | Product Location Char 4
Data 5 | Product Location Char 5
Data 6 | Product Location Char 6
Data 7 | Product Location Char 7
Data 8 | Product Location Char 8
Extended ID: | 0x101DXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x01D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 29
Frame: | 2 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Location Char 9
Data 2 | Product Location Char 10
Data 3 | Product Location Char 11
Data 4 | Product Location Char 12
Data 5 | Product Location Char 13
Data 6 | Product Location Char 14
Data 7 | Product Location Char 15
Data 8 | Product Location Char 16
Extended ID: | 0x101DXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x01D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 29
Frame: | 3 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Location Char 17
Data 2 | Product Location Char 18
Data 3 | Product Location Char 19
Data 4 | Product Location Char 20
Data 5 | Product Location Char 21
Data 6 | Product Location Char 22
Data 7 | Product Location Char 23
Data 8 | Product Location Char 24
Extended ID: | 0x101DXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x01D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x7
Function Code: | 29
Frame: | 4 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Product Location Char 25
Data 2 | Product Location Char 26
Data 3 | Product Location Char 27
Data 4 | Product Location Char 28
Data 5 | Product Location Char 29
Data 6 | Product Location Char 30
Data 7 | Product Location Char 31
FC 30: Sensor Data Configuration
Extended ID: | 0x101EXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x01E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 5
Function Code: | 30
Frame: | 1 of 1
Description: | Sensor Data Configuration- Configures what sensor data should be sent via CAN and how often.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Send interval (milliseconds) (Will be multiplied by 250 internally) (Bits 0-7)
Data 2 | Send interval (milliseconds) (Will be multiplied by 250 internally) (Bits 8-15)
Data 3 | Sensor ID ( Refer to 081-10015-03 for enumeration )
Data 4 | Config Type (Idle=0, In-Cycle=1, All=2)
Data 5 | #REF!
FC 31: Error History
Extended ID: | 0x101FXY85
Urgency: | 0x1 (Normal)
Function Code: | 0x01F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x0
DLC: | 0x8
Function Code: | 31
Frame: | 1 of 5
Description: | An array of errors from NVM
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 1 (Bits 0-7)
Data 2 | Error Code 1 (Bits 8-15)
Data 3 | Error Code 1 (Bits 16-23)
Data 4 | Error Code 1 (Bits 24-31)
Data 5 | Time/Date1 (Bit 0-7)
Data 6 | Time/Date1 (Bits 8-15)
Data 7 | Time/Date1 (Bits 16-23)
Data 8 | Time/Date1 (Bits 24-31)
Extended ID: | 0x101FXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x01F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 31
Frame: | 2 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 2 (Bits 0-7)
Data 2 | Error Code 2 (Bits 8-15)
Data 3 | Error Code 2 (Bits 16-23)
Data 4 | Error Code 2 (Bits 24-31)
Data 5 | Time/Date 2 (Bit 0-7)
Data 6 | Time/Date 2 (Bits 8-15)
Data 7 | Time/Date 2 (Bits 16-23)
Data 8 | Time/Date 2 (Bits 24-31)
Extended ID: | 0x101FXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x01F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 31
Frame: | 3 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 3 (Bits 0-7)
Data 2 | Error Code 3 (Bits 8-15)
Data 3 | Error Code 3 (Bits 16-23)
Data 4 | Error Code 3 (Bits 24-31)
Data 5 | Time/Date 3 (Bit 0-7)
Data 6 | Time/Date 3 (Bits 8-15)
Data 7 | Time/Date 3 (Bits 16-23)
Data 8 | Time/Date 3 (Bits 24-31)
Extended ID: | 0x101FXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x01F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 31
Frame: | 4 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 4 (Bits 0-7)
Data 2 | Error Code 4 (Bits 8-15)
Data 3 | Error Code 4 (Bits 16-23)
Data 4 | Error Code4 (Bits 24-31)
Data 5 | Time/Date 4 (Bit 0-7)
Data 6 | Time/Date 4 (Bits 8-15)
Data 7 | Time/Date 4 (Bits 16-23)
Data 8 | Time/Date 4 (Bits 24-31)
Extended ID: | 0x101FXYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x01F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 31
Frame: | 5 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Error Code 5 (Bits 0-7)
Data 2 | Error Code 5 (Bits 8-15)
Data 3 | Error Code 5 (Bits 16-23)
Data 4 | Error Code 5 (Bits 24-31)
Data 5 | Time/Date 5 (Bit 0-7)
Data 6 | Time/Date 5 (Bits 8-15)
Data 7 | Time/Date 5 (Bits 16-23)
Data 8 | Time/Date 5 (Bits 24-31)
FC 32: Diagnostic Mode
Extended ID: | 0x1020XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x020
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 32
Frame: | 1 of 1
Description: | Diagnostic Mode - Used to enter/exit diagnostic mode (Increased cloud data rates), may also be requested.
Origin | Display Interface / Connectivity module
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Command ( 0 = Diagnostic Mode Disabled, 1 = Diagnostic Mode Enabled)
Data 2 | #REF!
FC 35: Create Record
Extended ID: | 0x1023XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x023
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 35
Frame: | 1 of 1
Description: | Create Record - Sent by sterilizer control to notify the display a new cycle record is to be created
Origin | Sterilizer control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | Deprecated
FC 36: General Cycle Record (telemetry)
Extended ID: | 0x1024XY86
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Function Code: | 36
Frame: | 1 of 6
Description: | General Cycle Data - Duplicate of FC39, but sent by the display for telemetry to cloud
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Cycle Number per Day
Data 3 | Cycle Number (Bits 0-7)
Data 4 | Cycle Number (Bits 8-15)
Data 5 | #REF! | Unused | Unused | Unused | #REF!
Data 6 | Sterilizer Control Software Revision Major
Data 7 | Sterilizer Control Software Revision Minor
Data 8 | Sterilizer Control Software Revision Incidental
Extended ID: | 0x1024XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Frame: | 2 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Sterilizer Serial Number (Bits 0-7)
Data 2 | Sterilizer Serial Number (Bits 8-15)
Data 3 | Sterilizer Serial Number (Bits 16-23)
Data 4 | Sterilizer Serial Number (Bits 24-31)
Data 5 | Total Cycle Time in Seconds (Bits 0-7)
Data 6 | Total Cycle Time in Seconds (Bits 8-15)
Data 7 | Sterilization Temperature Setting tenths (Bits 0-7)
Data 8 | Sterilization Temperature Setting tenths (Bits 8-15)
Extended ID: | 0x1024XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Frame: | 3 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Sterilization Time Setting
Data 2 | Dry Time Setting
Data 3 | Cycle Start: Time/Date (Bits 0-7)
Data 4 | Cycle Start: Time/Date (Bits 8-15)
Data 5 | Cycle Start: Time/Date (Bits 16-23)
Data 6 | Cycle Start: Time/Date (Bits 24-31)
Data 7 | Cycle End: Time/Date (Bits 0-7)
Data 8 | Cycle End: Time/Date (Bits 8-15)
Extended ID: | 0x1024XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Frame: | 4 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle End: Time/Date (Bits 16-23)
Data 2 | Cycle End: Time/Date (Bits 24-31)
Data 3 | Cycle Fill Duration in Seconds (Bits 0-7)
Data 4 | Cycle Fill Duration in Seconds (Bits 8-15)
Data 5 | Cycle Heatup Duration in Seconds (Bits 0-7)
Data 6 | Cycle Heatup Duration in Seconds (Bits 8-15)
Data 7 | Cycle Sterilization Duration in Seconds (Bits 0-7)
Data 8 | Cycle Sterilization Duration in Seconds (Bits 8-15)
Extended ID: | 0x1024XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Frame: | 5 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Venting Duration in Seconds (Bits 0-7)
Data 2 | Cycle Venting Duration in Seconds (Bits 8-15)
Data 3 | Cycle Drying Duration in Seconds (Bits 0-7)
Data 4 | Cycle Drying Duration in Seconds (Bits 8-15)
Data 5 | Error Code (Bits 0-7)
Data 6 | Error Code (Bits 8-15)
Data 7 | Error Code (Bits 16-23)
Data 8 | Error Code (Bits 24-31)
Extended ID: | 0x1024XYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x024
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Frame: | 6 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Reserved |  | #REF! |  | #REF! |  | #REF!
FC 37: Pressure/Temperature Record (telemetry)
Extended ID: | 0x1025XY88
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x8
DLC: | 0x8
Function Code: | 37
Frame: | 1 of 8
Description: | Pressure/Temperature Samples - Duplicate of FC40, but sent by display for telemetry to cloud
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF! | Unused | Unused | Number of 15-Second Increments Between Samples
Data 4 | Number of samples in message |  |  |  | #REF!
Data 5 | Temperature in Tenths of Degree Sample 0 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 0 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 0 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 0 (Bits 8-15)
Extended ID: | 0x1025XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Frame: | 2 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 1 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 1 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 1 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 1 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 2 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 2 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 2 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 2 (Bits 8-15)
Extended ID: | 0x1025XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Frame: | 3 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 3 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 3 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 3 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 3 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 4 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 4 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 4 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 4 (Bits 8-15)
Extended ID: | 0x1025XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Frame: | 4 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 5 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 5 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 5 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 5 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 6 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 6 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 6 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 6 (Bits 8-15)
Extended ID: | 0x1025XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Frame: | 5 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 7 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 7 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 7 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 7 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 8 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 8 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 8 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 8 (Bits 8-15)
Extended ID: | 0x1025XYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Frame: | 6 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 9 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 9 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 9 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 9 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 10 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 10 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 10 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 10 (Bits 8-15)
Extended ID: | 0x1025XYC6
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Frame: | 7 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 11 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 11 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 11 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 11 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 12 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 12 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 12 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 12 (Bits 8-15)
Extended ID: | 0x1025XYC7
Urgency: | 0x1 (Normal)
Function Code: | 0x025
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x7
DLC: | 0x4
Frame: | 8 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 13 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 13 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 13 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 13 (Bits 8-15)
FC 38: Sterilization Summary (telemetry)
Extended ID: | 0x1026XY82
Urgency: | 0x1 (Normal)
Function Code: | 0x026
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 38
Frame: | 1 of 2
Description: | Sterilization Summary - Duplicate of FC41 but sent by display for telemetry to cloud
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF! | Sterilization Time in Minutes
Data 4 | Minimum Sterilization Temperature in Tenths of Degree (Bits 0-7)
Data 5 | Minimum Sterilization Temperature in Tenths of Degree (Bits 8-15)
Data 6 | Maximum Sterilization Temperature in Tenths of Degree (Bits 0-7)
Data 7 | Maximum Sterilization Temperature in Tenths of Degree (Bits 8-15)
Data 8 | Minimum Sterilization Pressure in Tenths of Unit (Bits 0-7)
Extended ID: | 0x1026XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x026
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x3
Frame: | 2 of 2
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Minimum Sterilization Pressure in Tenths of Unit (Bits 8-15)
Data 2 | Maximum Sterilization Pressure in Tenths of Unit (Bits 0-7)
Data 3 | Maximum Sterilization Pressure in Tenths of Unit (Bits 8-15)
FC 39: General Cycle Record
Extended ID: | 0x1027XY85
Urgency: | 0x1 (Normal)
Function Code: | 0x027
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Function Code: | 39
Frame: | 1 of 5
Description: | General Printout Data - To be used by the Display to provide general information about a previous cycle that has occurred.
Origin | Sterilizer Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Cycle Number per Day
Data 3 | Cycle Number (Bits 0-7)
Data 4 | Cycle Number (Bits 8-15)
Data 5 | #REF! | Unused | Unused | Unused | #REF!
Data 6 | Sterilizer Control Software Revision Major
Data 7 | Sterilizer Control Software Revision Minor
Data 8 | Sterilizer Control Software Revision Incidental
Extended ID: | 0x1027XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x027
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 39
Frame: | 2 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Sterilizer Serial Number (Bits 0-7)
Data 2 | Sterilizer Serial Number (Bits 8-15)
Data 3 | Sterilizer Serial Number (Bits 16-23)
Data 4 | Sterilizer Serial Number (Bits 24-31)
Data 5 | Total Cycle Time in Seconds (Bits 0-7)
Data 6 | Total Cycle Time in Seconds (Bits 8-15)
Data 7 | Sterilization Temperature Setting tenths (Bits 0-7)
Data 8 | Sterilization Temperature Setting tenths (Bits 8-15)
Extended ID: | 0x1027XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x027
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 39
Frame: | 3 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Sterilization Time Setting
Data 2 | Dry Time Setting
Data 3 | Cycle Start: Time/Date (Bits 0-7)
Data 4 | Cycle Start: Time/Date (Bits 8-15)
Data 5 | Cycle Start: Time/Date (Bits 16-23)
Data 6 | Cycle Start: Time/Date (Bits 24-31)
Data 7 | Cycle End: Time/Date (Bits 0-7)
Data 8 | Cycle End: Time/Date (Bits 8-15)
Extended ID: | 0x1027XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x027
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 39
Frame: | 4 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle End: Time/Date (Bits 16-23)
Data 2 | Cycle End: Time/Date (Bits 24-31)
Data 3 | Cycle Fill Duration in Seconds (Bits 0-7)
Data 4 | Cycle Fill Duration in Seconds (Bits 8-15)
Data 5 | Cycle Heatup Duration in Seconds (Bits 0-7)
Data 6 | Cycle Heatup Duration in Seconds (Bits 8-15)
Data 7 | Cycle Sterilization Duration in Seconds (Bits 0-7)
Data 8 | Cycle Sterilization Duration in Seconds (Bits 8-15)
Extended ID: | 0x1027XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x027
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 39
Frame: | 5 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Venting Duration in Seconds (Bits 0-7)
Data 2 | Cycle Venting Duration in Seconds (Bits 8-15)
Data 3 | Cycle Drying Duration in Seconds (Bits 0-7)
Data 4 | Cycle Drying Duration in Seconds (Bits 8-15)
Data 5 | Error Code (Bits 0-7)
Data 6 | Error Code (Bits 8-15)
Data 7 | Error Code (Bits 16-23)
Data 8 | Error Code (Bits 24-31)
FC 40: Pressure/Temperature Record
Extended ID: | 0x1028XY88
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x8
DLC: | 0x8
Function Code: | 40
Frame: | 1 of 8
Description: | Pressure/Temperature Samples - Sends samples of temperature and pressure data that was recorded during a specific cycle mode with a defined number of 15 second intervals between samples about a cycle that has occurred. It is expected for the Sterilizer Control to send this message out at the end of each applicable Cycle Mode (Heating, Sterilizing, Drying). The maximum amount of samples that can occur are 14 with each sample being a pair of two U16s. To prevent unnecessary complexity on modules that need to parse this, it should always be sent as an 8-frame message. If less than 14 samples were gathered, the unused bytes should be filled with 0xFF.
Origin | Sterilizer Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF! | Unused | Unused | Number of 15-Second Increments Between Samples
Data 4 | Number of samples in message |  |  |  | #REF!
Data 5 | Temperature in Tenths of Degree Sample 0 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 0 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 0 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 0 (Bits 8-15)
Extended ID: | 0x1028XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 40
Frame: | 2 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 1 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 1 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 1 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 1 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 2 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 2 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 2 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 2 (Bits 8-15)
Extended ID: | 0x1028XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 40
Frame: | 3 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 3 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 3 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 3 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 3 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 4 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 4 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 4 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 4 (Bits 8-15)
Extended ID: | 0x1028XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 40
Frame: | 4 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 5 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 5 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 5 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 5 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 6 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 6 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 6 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 6 (Bits 8-15)
Extended ID: | 0x1028XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 40
Frame: | 5 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 7 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 7 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 7 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 7 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 8 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 8 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 8 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 8 (Bits 8-15)
Extended ID: | 0x1028XYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Function Code: | 40
Frame: | 6 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 9 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 9 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 9 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 9 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 10 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 10 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 10 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 10 (Bits 8-15)
Extended ID: | 0x1028XYC6
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Function Code: | 40
Frame: | 7 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 11 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 11 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 11 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 11 (Bits 8-15)
Data 5 | Temperature in Tenths of Degree Sample 12 (Bits 0-7)
Data 6 | Temperature in Tenths of Degree Sample 12 (Bits 8-15)
Data 7 | Pressure in Tenths of Unit Sample 12 (Bits 0-7)
Data 8 | Pressure in Tenths of Unit Sample 12 (Bits 8-15)
Extended ID: | 0x1028XYC7
Urgency: | 0x1 (Normal)
Function Code: | 0x028
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x7
DLC: | 0x4
Function Code: | 40
Frame: | 8 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Temperature in Tenths of Degree Sample 13 (Bits 0-7)
Data 2 | Temperature in Tenths of Degree Sample 13 (Bits 8-15)
Data 3 | Pressure in Tenths of Unit Sample 13 (Bits 0-7)
Data 4 | Pressure in Tenths of Unit Sample 13 (Bits 8-15)
FC 41: Sterilization Summary
Extended ID: | 0x1029XY82
Urgency: | 0x1 (Normal)
Function Code: | 0x029
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 41
Frame: | 1 of 2
Description: | Sterilization Summary - Sends summary information about the sterilization cycle mode. It is expected for the Sterilizer Control to send this information when it completes the sterilization Cycle Mode.
Origin | Sterilizer Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF! | Sterilization Time in Minutes
Data 4 | Minimum Sterilization Temperature in Tenths of Degree (Bits 0-7)
Data 5 | Minimum Sterilization Temperature in Tenths of Degree (Bits 8-15)
Data 6 | Maximum Sterilization Temperature in Tenths of Degree (Bits 0-7)
Data 7 | Maximum Sterilization Temperature in Tenths of Degree (Bits 8-15)
Data 8 | Minimum Sterilization Pressure in Tenths of Unit (Bits 0-7)
Extended ID: | 0x1029XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x029
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x3
Function Code: | 41
Frame: | 2 of 2
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Minimum Sterilization Pressure in Tenths of Unit (Bits 8-15)
Data 2 | Maximum Sterilization Pressure in Tenths of Unit (Bits 0-7)
Data 3 | Maximum Sterilization Pressure in Tenths of Unit (Bits 8-15)
Data 4 | Calculated Altitude [feet] (Bits 0-7)
Data 5 | Calculated Altitude [feet] (Bits 8-15)
Data 6 | Pressure [mbar] (Bits 0-7)
Data 7 | Pressure [mbar] (Bits 8-15)
FC 42: Cycle Instruments/Packs
Extended ID: | 0x102AXY83
Urgency: | 0x1 (Normal)
Function Code: | 0x02A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 42
Frame: | 1 of 3
Description: | Cycle Instruments/Packs - Sends user-entered information about Instruments or Packs that are loaded in the chamber for a cycle including name and quantity of each. This multi-frame message will be sent once per instrument or pack for all instruments or packs for the cycle. It is expected for the Display to send this information when the cycle has completed.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | Quantity of Instrument / Pack
Data 4 | Instrument / Pack Name Char 0
Data 5 | Instrument / Pack Name Char 1
Data 6 | Instrument / Pack Name Char 2
Data 7 | Instrument / Pack Name Char 3
Data 8 | Instrument / Pack Name Char 4
Extended ID: | 0x102AXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x02A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 42
Frame: | 2 of 3
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Instrument / Pack Name Char 5
Data 2 | Instrument / Pack Name Char 6
Data 3 | Instrument / Pack Name Char 7
Data 4 | Instrument / Pack Name Char 8
Data 5 | Instrument / Pack Name Char 9
Data 6 | Instrument / Pack Name Char 10
Data 7 | Instrument / Pack Name Char 11
Data 8 | Instrument / Pack Name Char 12
Extended ID: | 0x102AXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x02A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x7
Function Code: | 42
Frame: | 3 of 3
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Instrument / Pack Name Char 13
Data 2 | Instrument / Pack Name Char 14
Data 3 | Instrument / Pack Name Char 15
Data 4 | Instrument / Pack Name Char 16
Data 5 | Instrument / Pack Name Char 17
Data 6 | Instrument / Pack Name Char 18
Data 7 | Record Type: Instrument (0) or Pack (1)
Data 8 | Instrument / Pack Name Char 18
Extended ID: | 0x102AXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x02A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 42
Frame: | 4 of 4
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Record Type: Instrument (0) or Pack (1)
FC 43: Cycle Indicator Information
Extended ID: | 0x102BXY87
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x7
DLC: | 0x8
Function Code: | 43
Frame: | 1 of 7
Description: | Cycle Indicator Information - Sends user-entered information about Chemical and Biological Indicators in a cycle. It is expected for the Display Interface to send any indicators information at the completion of each cycle.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF! |  | #REF! | #REF! | Unused |  | #REF!
Data 4 | Biological Indicator Lot Number Char 0
Data 5 | Biological Indicator Lot Number Char 1
Data 6 | Biological Indicator Lot Number Char 2
Data 7 | Biological Indicator Lot Number Char 3
Data 8 | Biological Indicator Lot Number Char 4
Extended ID: | 0x102BXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 43
Frame: | 2 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Biological Indicator Lot Number Char 5
Data 2 | Biological Indicator Lot Number Char 6
Data 3 | Biological Indicator Lot Number Char 7
Data 4 | Biological Indicator Lot Number Char 8
Data 5 | Biological Indicator Lot Number Char 9
Data 6 | Biological Indicator Expiration Day (1-31)
Data 7 | Biological Indicator Expiration Month (1-12)
Data 8 | Biological Indicator Expiration Year (2016 + X)
Extended ID: | 0x102BXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 43
Frame: | 3 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Time/Date Stamp (Bits 0-7)
Data 2 | Time/Date Stamp (Bits 8-15)
Data 3 | Time/Date Stamp (Bits 16-23)
Data 4 | Time/Date Stamp (Bits 24-31)
Data 5 | Biological Indicator Manufacturer Char 0
Data 6 | Biological Indicator Manufacturer Char 1
Data 7 | Biological Indicator Manufacturer Char 2
Data 8 | Biological Indicator Manufacturer Char 3
Extended ID: | 0x102BXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 43
Frame: | 4 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Biological Indicator Manufacturer Char 4
Data 2 | Biological Indicator Manufacturer Char 5
Data 3 | Biological Indicator Manufacturer Char 6
Data 4 | Biological Indicator Manufacturer Char 7
Data 5 | Biological Indicator Manufacturer Char 8
Data 6 | Biological Indicator Manufacturer Char 9
Data 7 | Biological Indicator Manufacturer Char 10
Data 8 | Biological Indicator Manufacturer Char 11
Extended ID: | 0x102BXYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 43
Frame: | 5 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Biological Indicator Manufacturer Char 12
Data 2 | Biological Indicator Manufacturer Char 13
Data 3 | Biological Indicator Manufacturer Char 14
Data 4 | Biological Indicator Manufacturer Char 15
Data 5 | Biological Indicator Manufacturer Char 16
Data 6 | Biological Indicator Manufacturer Char 17
Data 7 | Biological Indicator Manufacturer Char 18
Data 8 | Biological Indicator Manufacturer Char 19
Extended ID: | 0x102BXYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Function Code: | 43
Frame: | 6 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Biological Indicator Manufacturer Char 20
Data 2 | Biological Indicator Manufacturer Char 21
Data 3 | Biological Indicator Manufacturer Char 22
Data 4 | Biological Indicator Manufacturer Char 23
Data 5 | Biological Indicator Manufacturer Char 24
Data 6 | Biological Indicator Manufacturer Char 25
Data 7 | Biological Indicator Manufacturer Char 26
Data 8 | Biological Indicator Manufacturer Char 27
Extended ID: | 0x102BXYC6
Urgency: | 0x1 (Normal)
Function Code: | 0x02B
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x5
Function Code: | 43
Frame: | 7 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Biological Indicator Manufcaturer Char 28
Data 2 | Biological Indicator Manufcaturer Char 29
Data 3 | Biological Indicator Manufcaturer Char 30
Data 4 | Biological Indicator Manufcaturer Char 31
Data 5 | #REF!
FC 44: Cycle Notes
Extended ID: | 0x102CXY88
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x8
DLC: | 0x8
Function Code: | 44
Frame: | 1 of 8
Description: | Cycle Notes - Sends user-entered notes and associated timestamps. It is expected for the Display Interface to send cycle notes upon the end of the cycle (either Cylce Stop, an error, or after moving through all Cycle Completed screens). The User Authentication Index is set by the display at the time the note is created to correlate the user authentication associated with this specific note record.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | Time/Date Stamp (Bits 0-7)
Data 4 | Time/Date Stamp (Bits 8-15)
Data 5 | Time/Date Stamp (Bits 16-23)
Data 6 | Time/Date Stamp (Bits 24-31)
Data 7 | Authentication Reason 0 = Cycle Start, 1 = Cycle Complete, 2 = Cycle Interruption, 3 = Cycle: Biological Indicator Result, 9 = Cycle Start Note 1, 10 = Cycle Start Note 2, 11 = Cycle Release Note 1, 12 = Cycle Release Note 2, 13 = BI Note, 14 = Cycle History Note 1, 15 = Cycle History Note 2
Data 8 | Note Char 0
Extended ID: | 0x102CXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 44
Frame: | 2 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 1
Data 2 | Note Char 2
Data 3 | Note Char 3
Data 4 | Note Char 4
Data 5 | Note Char 5
Data 6 | Note Char 6
Data 7 | Note Char 7
Data 8 | Note Char 8
Extended ID: | 0x102CXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 44
Frame: | 3 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 9
Data 2 | Note Char 10
Data 3 | Note Char 11
Data 4 | Note Char 12
Data 5 | Note Char 13
Data 6 | Note Char 14
Data 7 | Note Char 15
Data 8 | Note Char 16
Extended ID: | 0x102CXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 44
Frame: | 4 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 17
Data 2 | Note Char 18
Data 3 | Note Char 19
Data 4 | Note Char 20
Data 5 | Note Char 21
Data 6 | Note Char 22
Data 7 | Note Char 23
Data 8 | Note Char 24
Extended ID: | 0x102CXYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 44
Frame:
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 25
Data 2 | Note Char 26
Data 3 | Note Char 27
Data 4 | Note Char 28
Data 5 | Note Char 29
Data 6 | Note Char 30
Data 7 | Note Char 31
Data 8 | Note Char 32
Extended ID: | 0x102CXYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x8
Function Code: | 44
Frame: | 6 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 33
Data 2 | Note Char 34
Data 3 | Note Char 35
Data 4 | Note Char 36
Data 5 | Note Char 37
Data 6 | Note Char 38
Data 7 | Note Char 39
Data 8 | Note Char 40
Extended ID: | 0x102CXYC6
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Function Code: | 44
Frame: | 7 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 41
Data 2 | Note Char 42
Data 3 | Note Char 43
Data 4 | Note Char 44
Data 5 | Note Char 45
Data 6 | Note Char 46
Data 7 | Note Char 47
Data 8 | Note Char 48
Extended ID: | 0x102CXYC7
Urgency: | 0x1 (Normal)
Function Code: | 0x02C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x7
DLC: | 0x3
Function Code: | 44
Frame: | 8 of 8
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Note Char 49
Data 2 | Note Char 50
Data 3 | Note Char 51
FC 46: BI Results
Extended ID: | 0x102EXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x02E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x8
Function Code: | 46
Frame: | 1 of 2
Description: | BI Results - Sends user-entered information about Biological Indicator results and a current timestamp.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | Unused | [5 empty] | #REF!
Data 4 | Time/Date Stamp (Bits 0-7)
Data 5 | Time/Date Stamp (Bits 8-15)
Data 6 | Time/Date Stamp (Bits 16-23)
Data 7 | Time/Date Stamp (Bits 24-31)
Data 8 | #REF!
FC 50: Routine Care Record
Extended ID: | 0x1032XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x032
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x7
Function Code: | 50
Frame: | 1 of 1
Description: | Routine Care Record - A message reporting Routine Care record information. This is broadcast whenever a User completes a Routine Care action on the Display. This includes Part Replacement actions.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | RCM Record Number (Bits 0 - 7)
Data 2 | RCM Record Number (Bits 8 - 15)
Data 3 | Time/Date Stamp (Bits 0 - 7)
Data 4 | Time/Date Stamp (Bits 8 - 15)
Data 5 | Time/Date Stamp (Bits 16 - 23)
Data 6 | Time/Date Stamp (Bits 24 - 31)
Data 7 | #REF!
Data 8 | Rinsing Cycle | Cleaning Cycle | Checked Filters | Cleaned Chamber | Cleaned Trays | Drained Reservoir | Inspected Gaskets | Cleaned Surfaces
Extended ID: | 0x1032XYC1
Function Code: | 0x039
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x1
Frame: | 2 of 2
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Reserved | Reserved | Reserved | Replaced Door Gasket | Replaced Dam Gasket | Wiped Chamber | Wiped Trays | Tested PRV
FC 53: Language/Display Units
Extended ID: | 0x1035XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x035
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3
Function Code: | 53
Frame: | 1 of 1
Description: | Language/Display Units - The language for the display text (English, Spanish or French) and the units (Standard or Metric) for the temperature and pressure (F/PSI or C/kPa).
Origin | Display, Sterilization Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | #REF!
Data 3 | #REF!
FC 57: Custom Cycle Name
Extended ID: | 0x1039XY83
Urgency: | 0x1 (Normal)
Function Code: | 0x039
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 57
Frame: | 1 of 3
Description: | Custom Cycle Name - The user-entered name for the various custom cycles
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF!
Data 4 | Custom Cycle Name Char 0
Data 5 | Custom Cycle Name Char 1
Data 6 | Custom Cycle Name Char 2
Data 7 | Custom Cycle Name Char 3
Data 8 | Custom Cycle Name Char 4
Extended ID: | 0x1039XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x039
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 57
Frame: | 2 of 3
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Custom Cycle Name Char 5
Data 2 | Custom Cycle Name Char 6
Data 3 | Custom Cycle Name Char 7
Data 4 | Custom Cycle Name Char 8
Data 5 | Custom Cycle Name Char 9
Data 6 | Custom Cycle Name Char 10
Data 7 | Custom Cycle Name Char 11
Data 8 | Custom Cycle Name Char 12
Extended ID: | 0x1039XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x039
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x2
Function Code: | 57
Frame: | 3 of 3
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Custom Cycle Name Char 13
Data 2 | Custom Cycle Name Char 14
FC 58: Bowie-Dick Test Pack Information
Extended ID: | 0x103AXY87
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x7
DLC: | 0x8
Function Code: | 58
Frame: | 1 of 7
Description: | Bowie-Dick Test Pack Information - Sends user-entered information about the Bowie-Dick Test Pack in a Bowie-Dick cycle. It is expected for the Display Interface to send any indicators information at the completion of each cycle.
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | Unused | Unused | Unused | Unused | Unused | Unused | #REF!
Data 4 | Test Pack Lot Number Char 0
Data 5 | Test Pack Lot Number Char 1
Data 6 | Test Pack Lot Number Char 2
Data 7 | Test Pack Lot Number Char 3
Data 8 | Test Pack Lot Number Char 4
Extended ID: | 0x103AXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 58
Frame: | 2 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Test Pack Lot Number Char 5
Data 2 | Test Pack Lot Number Char 6
Data 3 | Test Pack Lot Number Char 7
Data 4 | Test Pack Lot Number Char 8
Data 5 | Test Pack Lot Number Char 9
Data 6 | Test Pack Expiration Day (1-31)
Data 7 | Test Pack Expiration Month (1-12)
Data 8 | Test Pack Expiration Year (2016 + X)
Extended ID: | 0x103AXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 58
Frame: | 3 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Time/Date Stamp (Bits 0-7)
Data 2 | Time/Date Stamp (Bits 8-15)
Data 3 | Time/Date Stamp (Bits 16-23)
Data 4 | Time/Date Stamp (Bits 24-31)
Data 5 | Test Pack Manufacturer Char 0
Data 6 | Test Pack Manufacturer Char 1
Data 7 | Test Pack Manufacturer Char 2
Data 8 | Test Pack Manufacturer Char 3
Extended ID: | 0x103AXYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 58
Frame: | 4 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Test Pack Manufacturer Char 4
Data 2 | Test Pack Manufacturer Char 5
Data 3 | Test Pack Manufacturer Char 6
Data 4 | Test Pack Manufacturer Char 7
Data 5 | Test Pack Manufacturer Char 8
Data 6 | Test Pack Manufacturer Char 9
Data 7 | Test Pack Manufacturer Char 10
Data 8 | Test Pack Manufacturer Char 11
Extended ID: | 0x103AXYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 58
Frame: | 5 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Test Pack Manufacturer Char 12
Data 2 | Test Pack Manufacturer Char 13
Data 3 | Test Pack Manufacturer Char 14
Data 4 | Test Pack Manufacturer Char 15
Data 5 | Test Pack Manufacturer Char 16
Data 6 | Test Pack Manufacturer Char 17
Data 7 | Test Pack Manufacturer Char 18
Data 8 | Test Pack Manufacturer Char 19
Extended ID: | 0x103AXYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 58
Frame: | 6 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Test Pack Manufacturer Char 20
Data 2 | Test Pack Manufacturer Char 21
Data 3 | Test Pack Manufacturer Char 22
Data 4 | Test Pack Manufacturer Char 23
Data 5 | Test Pack Manufacturer Char 24
Data 6 | Test Pack Manufacturer Char 25
Data 7 | Test Pack Manufacturer Char 26
Data 8 | Test Pack Manufacturer Char 27
Extended ID: | 0x103AXYC6
Urgency: | 0x1 (Normal)
Function Code: | 0x03A
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x4
Function Code: | 58
Frame: | 7 of 7
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Test Pack Manufacturer Char 28
Data 2 | Test Pack Manufacturer Char 29
Data 3 | Test Pack Manufacturer Char 30
Data 4 | Test Pack Manufacturer Char 31
FC 60: Sensor Control Component Status
Extended ID: | 0x103CXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x03C
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 60
Frame: | 1 of 1
Description: | Sensor Control Component status - A message for conveying the status of the sensor Control components.
Origin | Sensor Control
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Unused | Unused | #REF! | #REF! | #REF! | #REF! | #REF! | Unused
FC 61: Sensor Control Component Control
Extended ID: | 0x103DXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x03D
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 61
Frame: | 1 of 1
Description: | Sensor Control Component Control - Used as a command to turn on/off a single component on the sensor Control.
Origin | Varies upon sending node
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | State (0=Off, 1=On)
FC 62: Diagnostic Component Control
Extended ID: | 0x103EXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x03E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 62
Frame: | 1 of 1
Description: | Diagnostic Component Control - Used as a command to turn on/off a single sterilization component for diagnostic purposes.
Origin | Varies upon sending node
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Not Used | Not Used | Door Open Motor | Heater Relay | Fill Relay | Vent Relay | Air Relay | Not Used
Data 2 | State (0=Off, 1=On)
FC 63: Auto Fill Request
Extended ID: | 0x103FXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x03F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 63
Frame: | 1 of 1
Description: | Auto Fill Request - Used as a request turn on the auto fill valve when the Fill mode time indicates a low level of water in the reservoir.
Origin | Sterilization Control node
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | State (0=Off, 1=On)
FC 70: User Authentication - First Name
Extended ID: | 0x1046XY86
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Function Code: | 70
Frame: | 1 of 6
Description: | User Authentication - Sends information about when a user authenticates in the display. Authentications can happen for a variety of reasons and may need to be tracked later (Ex. "Cycle": In Cycle History, "Maintenance": In Maintenance History). This tracking may need to include timestamps, cycle count, or maintenance history record count, and should always contain the user name.
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF!
Data 4 | Time/Date Stamp (Bits 0-7)
Data 5 | Time/Date Stamp (Bits 8-15)
Data 6 | Time/Date Stamp (Bits 16-23)
Data 7 | Time/Date Stamp (Bits 24-31)
Data 8 | User First Name Char 0
Extended ID: | 0x1046XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 70
Frame: | 2 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User First Name Char 1
Data 2 | User First Name Char 2
Data 3 | User First Name Char 3
Data 4 | User First Name Char 4
Data 5 | User First Name Char 5
Data 6 | User First Name Char 6
Data 7 | User First Name Char 7
Data 8 | User First Name Char 8
Extended ID: | 0x1046XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 70
Frame: | 3 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User First Name Char 9
Data 2 | User First Name Char 10
Data 3 | User First Name Char 11
Data 4 | User First Name Char 12
Data 5 | User First Name Char 13
Data 6 | User First Name Char 14
Data 7 | User First Name Char 15
Data 8 | User First Name Char 16
Extended ID: | 0x1046XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 70
Frame: | 4 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User First Name Char 17
Data 2 | User First Name Char 18
Data 3 | User First Name Char 19
Data 4 | User First Name Char 20
Data 5 | User First Name Char 21
Data 6 | User First Name Char 22
Data 7 | User First Name Char 23
Data 8 | User First Name Char 24
Extended ID: | 0x1046XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 70
Frame: | 5 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User First Name Char 25
Data 2 | User First Name Char 26
Data 3 | User First Name Char 27
Data 4 | User First Name Char 28
Data 5 | User First Name Char 29
Data 6 | User First Name Char 30
Data 7 | User First Name Char 31
Data 8 | RCM Number (Bits 0-7)
Extended ID: | 0x1046XYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x046
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x1
Function Code: | 70
Frame: | 6 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | RCM Number (Bits 8-15)
FC 71: User Authentication - Last Name
Extended ID: | 0x1047XY86
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x6
DLC: | 0x8
Function Code: | 71
Frame: | 1 of 6
Description: | User Authentication - Sends information about when a user authenticates in the display. Authentications can happen for a variety of reasons and may need to be tracked later (Ex. "Cycle": In Cycle History, "Maintenance": In Maintenance History). This tracking may need to include timestamps, cycle count, or maintenance history record count, and should always contain the user name.
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Cycle Number (Bits 0-7)
Data 2 | Cycle Number (Bits 8-15)
Data 3 | #REF!
Data 4 | Time/Date Stamp (Bits 0-7)
Data 5 | Time/Date Stamp (Bits 8-15)
Data 6 | Time/Date Stamp (Bits 16-23)
Data 7 | Time/Date Stamp (Bits 24-31)
Data 8 | User Last Name Char 0
Extended ID: | 0x1047XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 71
Frame: | 2 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Last Name Char 1
Data 2 | User Last Name Char 2
Data 3 | User Last Name Char 3
Data 4 | User Last Name Char 4
Data 5 | User Last Name Char 5
Data 6 | User Last Name Char 6
Data 7 | User Last Name Char 7
Data 8 | User Last Name Char 8
Extended ID: | 0x1047XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 71
Frame: | 3 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Last Name Char 9
Data 2 | User Last Name Char 10
Data 3 | User Last Name Char 11
Data 4 | User Last Name Char 12
Data 5 | User Last Name Char 13
Data 6 | User Last Name Char 14
Data 7 | User Last Name Char 15
Data 8 | User Last Name Char 16
Extended ID: | 0x1047XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 71
Frame: | 4 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Last Name Char 17
Data 2 | User Last Name Char 18
Data 3 | User Last Name Char 19
Data 4 | User Last Name Char 20
Data 5 | User Last Name Char 21
Data 6 | User Last Name Char 22
Data 7 | User Last Name Char 23
Data 8 | User Last Name Char 24
Extended ID: | 0x1047XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 71
Frame: | 5 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Last Name Char 25
Data 2 | User Last Name Char 26
Data 3 | User Last Name Char 27
Data 4 | User Last Name Char 28
Data 5 | User Last Name Char 29
Data 6 | User Last Name Char 30
Data 7 | User Last Name Char 31
Data 8 | RCM Number (Bits 0-7)
Extended ID: | 0x1047XYC5
Urgency: | 0x1 (Normal)
Function Code: | 0x047
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x5
DLC: | 0x1
Function Code: | 71
Frame: | 6 of 6
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | RCM Number (Bits 8-15)
FC 78: BI Usage Notification
Extended ID: | 0x104EXY00
Urgency: | 0x1 (Normal)
Function Code: | 0x04E
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 78
Frame: | 1 of 1
Description: | BI Usage Notification - Sent to cloud to indicate due date for using a BI. 0 = Due today. Negative number = overdue. Positive number = Due in the future
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Days from Due (Bits 0-7) (signed)
Data 2 | Days from Due (Bits 8-15) (signed)
FC 79: RCM Notification
Extended ID: | 0x1023XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x04F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x7
Function Code: | 79
Frame: | 1 of 1
Description: | RCM Notification - Sent to cloud when RCM notifications are displayed on the product
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Last Performed Date/Time (Bits 0-7)
Data 3 | Last Performed Date/Time (Bits 8-15)
Data 4 | Last Performed Date/Time (Bits 16-23)
Data 5 | Last Performed Date/Time (Bits 24-31)
Data 6 | Days from Due (Bits 0-7) (signed)
Data 7 | Days from Due (Bits 8-15) (signed)
FC 80: Other Notifications
Extended ID: | 0x1050XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x050
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 80
Frame: | 1 of 1
Description: | RCM Additional Notification - Sent to cloud when RCM notifications are displayed on the product
Origin | Display Interface
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
FC 81: Component Measurements
Urgency: | 0x1 (Normal)
Function Code: | 0x051
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 81
Frame: | 1 of 2
Description: | Component Measurements: This frame is specifically used for diagnostic functionality related to Component Check and System Diagnostics features on the Display. If the Display were to want to start a component measurement, it would be done by sending the Sensor Type of the chosen component with all other data as Unused (0xFF). When the Sensor Control wants to report back the results of the component measurement, it would send the Sensor Type of the chosen component with all data filled in. The High Calibration and Low Calibration values are used as bounding values for the normal range of operation. If the Measured Value is above the High Calibration Value, an overcurrent condition results. If the Measured Value is below the Low Calibration Value, an undercurrent condition results. If the High Calibration and Low Calibration values are to be set, the Sensor Type of the chosen component should be used, the Measured Value data should be sent as Unused (0xFF). These calibration values could be sent by the Display Interface (if inserted via USB), the Connectivity Module (if sent via cloud), or the Test Interface (at the end of line test station).
Origin | Sensor Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Measured Value (Float; Bits 0-7)
Data 3 | Measured Value (Float; Bits 8-15)
Data 4 | Measured Value (Float; Bits 16-23)
Data 5 | Measured Value (Float; Bits 24-31)
Data 6 | High Calibration Value (Float; Bits 0-7)
Data 7 | High Calibration Value (Float; Bits 8-15)
Data 8 | High Calibration Value (Float; Bits 16-23
Extended ID: | 0x1051XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x051
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x5
Function Code: | 81
Frame: | 2 of 2
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | High Calibration Value (Float; Bits 24-31)
Data 2 | Low Calibration Value (Float; Bits 0-7)
Data 3 | Low Calibration Value (Float; Bits 8-15)
Data 4 | Low Calibration Value (Float; Bits 16-23)
Data 5 | Low Calibration Value (Float; Bits 24-31)
FC 82: Altimeter Information
Extended ID: | 0x1052XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x052
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x6
Function Code: | 82
Frame: | 1 of 1
Description: | Altimeter Information - This function code provides the current pressure [mbars], temperature [deg C] reading and calculated altitude [feet] from the altimeter connected to the Sterilizer Control. If desired, it should be requested through a Function Code 0: Message Request.
Origin | Sterilizer Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Calculated Altitude [feet] (Bits 0-7)
Data 2 | Calculated Altitude [feet] (Bits 8-15)
Data 3 | Pressure [mbar] (Bits 0-7)
Data 4 | Pressure [mbar] (Bits 8-15)
Data 5 | Temperature - [deg C] (Bits 0-7)
Data 6 | Temperature [deg C] (Bits 8-15)
FC 84: Constant Voltage references
Extended ID: | 0x1054XY82
Urgency: | 0x1 (Normal)
Function Code: | 0x054
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 84
Frame: | 1 of 2
Description: | Constant Voltage references - Contains the floating point values of the constant voltage references on the Sterilizer control PCBA
Origin | Sensor Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Fixed 3.0V reference voltage (bits 0-7)
Data 2 | Fixed 3.0V reference voltage (bits 8-15)
Data 3 | Fixed 3.0V reference voltage (bits 16-23)
Data 4 | Fixed 3.0V reference voltage (bits 24-31)
Data 5 | Fixed 1.5V reference voltage (bits 0-7)
Data 6 | Fixed 1.5V reference voltage (bits 8-15)
Data 7 | Fixed 1.5V reference voltage (bits 16-23)
Data 8 | Fixed 1.5V reference voltage (bits 24-31)
Extended ID: | 0x1054XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x054
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (continuous frame)
Frame Count: | 0x1
DLC: | 0x4
Function Code: | 84
Frame: | 2 of 2
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Fixed 2.5V reference voltage (bits 0-7)
Data 2 | Fixed 2.5V reference voltage (bits 8-15)
Data 3 | Fixed 2.5V reference voltage (bits 16-23)
Data 4 | Fixed 2.5V reference voltage (bits 24-31)
FC 85: Voltage Readings
Extended ID: | 0x1055XY82
Urgency: | 0x1 (Normal)
Function Code: | 0x055
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 85
Frame: | 1 of 2
Description: | Voltage Readings - Contains telemetry readings for the 12V divider, 5V divider, and AC Voltage
Origin | Sensor Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | AC voltage (bits 0-7)
Data 2 | AC voltage (bits 8-15)
Data 3 | AC voltage (bits 16-23)
Data 4 | AC voltage (bits 24-31)
Data 5 | 12V Rail voltage (bits 0-7)
Data 6 | 12V Rail voltage (bits 8-15)
Data 7 | 12V Rail voltage (bits 16-23)
Data 8 | 12V Rail voltage (bits 24-31)
Extended ID: | 0x1055XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x055
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (continuous frame)
Frame Count: | 0x1
DLC: | 0x4
Function Code: | 85
Frame: | 2 of 2
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | 5V Rail voltage (bits 0-7)
Data 2 | 5V Rail voltage (bits 8-15)
Data 3 | 5V Rail voltage (bits 16-23)
Data 4 | 5V Rail voltage (bits 24-31)
FC 86: Current Readings
Extended ID: | 0x1056XY82
Urgency: | 0x1 (Normal)
Function Code: | 0x056
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 86
Frame: | 1 of 2
Description: | Current Readings - Contains telemetry readings for the AC current and 12V DC current
Origin | Sensor Control
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | AC Current 1 (bits 0-7)
Data 2 | AC Current 1 (bits 8-15)
Data 3 | AC Current 1 (bits 16-23)
Data 4 | AC Current 1 (bits 24-31)
Data 5 | AC Current 2 (bits 0-7)
Data 6 | AC Current 2 (bits 8-15)
Data 7 | AC Current 2 (bits 16-23)
Data 8 | AC Current 2 (bits 24-31)
Extended ID: | 0x1056XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x056
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (continuous frame)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 86
Frame: | 2 of 2
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | AC Current 3 (bits 0-7)
Data 2 | AC Current 3 (bits 8-15)
Data 3 | AC Current 3 (bits 16-23)
Data 4 | AC Current 3 (bits 24-31)
Data 5 | 12V DC rail Current (bits 0-7)
Data 6 | 12V DC rail Current (bits 8-15)
Data 7 | 12V DC rail Current (bits 16-23)
Data 8 | 12V DC rail Current (bits 24-31)
FC 88: RCM Event
Extended ID: | 0x1058XY86
Urgency: | 0x1 (Normal)
Function Code: | 0x058
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 88
Frame: | 1 of 2
Description: | General Cycle Data - Duplicate of FC39, but sent by the display for telemetry to cloud
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | RCM Number (Bits 0-7)
Data 3 | RCM Number (Bits 8-15)
Data 4 | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF! | #REF!
Data 5 | Unused |  |  | #REF! | #REF! | #REF! | #REF! | #REF!
Data 6 | Time/Date (Bit 0-7)
Data 7 | Time/Date (Bits 8-15)
Data 8 | Time/Date (Bits 16-23)
Extended ID: | 0x1058XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x058
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x1
Function Code: | 88
Frame: | 2 of 2
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Time/Date (Bits 24-31)
FC 89: RCM Status
Extended ID: | 0x1059XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x059
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x6
Function Code: | 89
Frame: | 1 of 1
Description: | RCM Status - This message is sent whenever an RCM event happens or when numbers count down every day.
Origin | Display
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Daily Days From Due (Bits 0-7)
Data 2 | Daily Days From Due (Bits 8-15)
Data 3 | Weekly Days From Due (Bits 0-7)
Data 4 | Weekly Days From Due (Bits 8-15)
Data 5 | Monthly Days From Due (Bits 0-7)
Data 6 | Monthly Days From Due (Bits 8-15)
FC 98: Product Operational Status
Extended ID: | 0x1062XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x062
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 98
Frame: | 1 of 1
Description: | Operational status - Sent by the display for telemetry purposes every 5 minutes
Origin | Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | #REF!
FC 99: Product In-Use Status
Extended ID: | 0x1063XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x063
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 99
Frame: | 1 of 1
Description: | Product In-Use Status - A message reporting whether the product is "In-use"
Origin | Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | In Use Status (Not in Use = 0, In Use = 1)
FC 100: Device Settings (Single Frame)
Extended ID: | 0x1064XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x064
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | (Varies based on setting length)
Function Code: | 100
Frame: | 1 of 1
Description: | Single frame message for device settings. Device settings are meant to be sent from the cloud to the device. DLC is dependant on the "Value Type". Refer to the Device settings sheet in this document.
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Setting ID (Bits 0-7) (Refer to Device Settings Sheet)
Data 2 | Setting ID (Bits 8-15) (Refer to Device Settings Sheet)
Data 3 | #REF! | #REF!
Data 4 | RESERVED |  |  |  | #REF!
Data 5 | Data (Bits 0-7) (Mandatory)
Data 6 | Data (Bits 8-15) (Optional)
Data 7 | Data (Bits 16-23) (Optional)
Data 8 | Data (Bits 24-31) (Optional)
FC 101: Device Settings Ack/Nack
Extended ID: | 0x1065XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x065
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 101
Frame: | 1 of 1
Description: | Message sent to the connectivity module as a response when a device setting is received.
Origin | Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Setting ID (Bits 0-7) (Refer to Device Settings Sheet)
Data 2 | Setting ID (Bits 8-15) (Refer to Device Settings Sheet)
Data 3 | #REF!
FC 102: Device Settings (Multi Frame)
Extended ID: | 0x1066XY8Z
Urgency: | 0x1 (Normal)
Function Code: | 0x066
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0xZ (Varies based on setting length)
DLC: | 0x08
Function Code: | 102
Frame: | 1 of Z
Description: | Multi frame message for device settings. Device settings are meant to be sent from the cloud to the device. DLC and number of frames are dependant on the "Value Type". Refer to the Device settings sheet in this document.
Origin | Any
Frequency | Connectivity Module
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | Setting ID (Bits 0-7) (Refer to Device Settings Sheet)
Data 2 | Setting ID (Bits 8-15) (Refer to Device Settings Sheet)
Data 3 | #REF! | #REF!
Data 4 | RESERVED |  |  |  | #REF!
Data 5 | Data (Bits 0-7) (Mandatory)
Data 6 | Data (Bits 8-15) (Optional)
Data 7 | Data (Bits 16-23) (Optional)
Optional Data 8 | Data (Bits 24-31) (Optional)
Extended ID: | 0x1066XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x066
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | (Varies based on setting length)
Function Code: | 102
Frame: | 2 of Z
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Data Continued (Optional)
Data 2 | Data Continued (Optional)
Data 3 | Data Continued (Optional)
Data 4 | Data Continued (Optional)
Data 5 | Data Continued (Optional)
Data 6 | Data Continued (Optional)
Data 7 | Data Continued (Optional)
Optional Data 8 | Data Continued (Optional)
More frames optional for strings over 11 characters in length
FC 103: Device Settings Control
Extended ID: | 0x1067XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x067
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x4
Function Code: | 103
Frame: | 1 of 1
Description: | Device Settings Control - Used to signal the start or end of device settings process or device settings bucket
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | #REF!
Data 2 | Reserved |  |  |  | #REF!
Data 3 | Number of Settings (Bits 0-7)
Data 4 | Number of Settings (Bits 8-15)
FC 104: Device Settings Version
Extended ID: | 0x1068XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x068
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x1
Function Code: | 104
Frame: | 1 of 1
Description: | Device Settings Version - May be requested to check what version of the device settings API a node complies with
Origin | Any
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Device settings API Version (Version 1=1, Version 2=2, etc.)
FC 105: Device Settings Request
Extended ID: | 0x1069XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x069
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 105
Frame: | 1 of 1
Description: | Device Settings Request - Used to request a device setting by ID.
Origin | Any
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Setting ID (Bits 0-7) (Refer to Device Settings Sheet)
Data 2 | Setting ID (Bits 8-15) (Refer to Device Settings Sheet)
FC 128: Hardware/Software Information Metadata
Extended ID: | 0x1080XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x080
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 1
Function Code: | 128
Frame: | 1 of 1
Description: | Precursor message sent prior to send Hw/Sw Info. Tells any receiving nodes how many messages to expect in response to a request for hardware/software info. When this message is requested, it will be returned along with Hw/Sw Info. After requesting this message, the requester shall allow up to 1 second for response messages to be sent.
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Number of Hw/Sw Information messages to expect
FC 129-159: Hardware/Software Information
Extended ID: | 0x1081XY83 - 0x109FXY83
Urgency: | 0x1 (Normal)
Function Code: | 0x81 - 0x9F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 3
DLC: | 0x08
Function Code: | 129-159
Frame: | 1 of 3
Description: | The software and hardware part/version information. This message may need to be sent multiple times in succession by a CAN node if it has multiple pieces of information. In this case, the first message to be sent shall be function code 0x81 and each message after that will increment the function code. A node may also send this message on behalf of downstream hardware/software that it speaks for on the CAN bus. The function codes that each node ID uses for each piece of hardware/software information are product specific and shall be documented with other product specific CAN data. The specific function code associated with a particular piece of information on a node is known as the “Information Element” number. Refer to HW/SW info sheet of this document for the definitions of information element number for each node. Hw/Sw Info Metada should be sent prior to this message. Hw/Sw Info metadata is to be requested rather than this message.
Origin | Any
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | (0=Updatable, 1=Not Updatable) | Type ID (0=Midmark)
Data 2 | Software Part Number (Bits 0-7)
Data 3 | Software Part Number (Bits 8-15)
Data 4 | Software Part Number (Bits 16-23)
Data 5 | Software Part Number (Bits 24-13)
Data 6 | Software Version Major
Data 7 | Software Version Minor
Data 8 | Software Version Incidental
Extended ID: | 0x1081XYC1 - 0x109FXYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x81 - 0x9F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 8
Function Code: | 129-159
Frame: | 2 of 3
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Number of Pages |  |  |  | Active Page (1 based)
Data 2 | PCB Serial Number (Bits 0-7)
Data 3 | PCB Serial Number (Bits 8-15)
Data 4 | PCB Serial Number (Bits 16-23)
Data 5 | PCB Serial Number (Bits 24-31)
Data 6 | PCB Assembly Item Number (Bits 0-7)
Data 7 | PCB Assembly Item Number (Bits 8-15)
Optional Data 8 | PCB Assembly Item Number (Bits 16-23)
Extended ID: | 0x1081XYC2 - 0x109FXYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x81 - 0x9F
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 4
Function Code: | 129-159
Frame: | 3 of 3
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | PCB Assembly Item Number (Bits 24-31)
Data 2 | PCB Assembly Revision (left most letter in ASCII)
Data 3 | PCB Assembly Revision (second letter in ASCII)
Data 4 | PCB Assembly Revision (number in ASCII)
FC 212: User Accounts
Extended ID: | 0x10D4XY85
Urgency: | 0x1 (Normal)
Function Code: | 0x0D4
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x2 (First Frame of Multi-Frame Message)
Frame Count: | 0x05
DLC: | 0x8
Function Code: | 212
Frame: | 1 of 5
Description: | User Accounts: Users are able to make accounts within the NGS2 Display Interface or through the Midmark portal. These accounts are to be used for User Authentication functionality, which is a digital signature of the user performing some action on the device. Since users are able to be created, modified, or deleted on the NGS2 Display Interface or the Midmark Portal (relayed to this product via the Connectivity Module), this multi-frame message must exist to allow synchronization between the two systems. Since the user fields are longer than what can fit in a single multi-frame message, it is expected that any edits to user accounts take place in a batched fashion, where the sender can group multiple multi-frame messages together by identifying them as "User Account 1 of X", "User Account 2 of X", etc. Each User Account Field will be sent as up to 32 ASCII characters, including PINs which shall be only 4 numeric characters. Unused characters should be sent as 0x00.
Origin | Display Interface, Connectivity Module
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | #REF!
Data 2 | Total User Account Actions to Occur (X Actions)
Data 3 | Current User Account Action Index (1 of X Actions, 2 of X Actions, etc.)
Data 4 | Time/Date of User Account Action (Bits 0-7)
Data 5 | Time/Date of User Account Action (Bits 8-15)
Data 6 | Time/Date of User Account Action (Bits 16-23)
Data 7 | Time/Date of User Account Action (Bits 24-31)
Data 8 | #REF!
Extended ID: | 0x10D4XYC1
Urgency: | 0x1 (Normal)
Function Code: | 0x0D4
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x1
DLC: | 0x8
Function Code: | 212
Frame: | 2 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Account Field Char 0
Data 2 | User Account Field Char 1
Data 3 | User Account Field Char 2
Data 4 | User Account Field Char 3
Data 5 | User Account Field Char 4
Data 6 | User Account Field Char 5
Data 7 | User Account Field Char 6
Data 8 | User Account Field Char 7
Extended ID: | 0x10D4XYC2
Urgency: | 0x1 (Normal)
Function Code: | 0x0D4
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x2
DLC: | 0x8
Function Code: | 212
Frame: | 3 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Account Field Char 8
Data 2 | User Account Field Char 9
Data 3 | User Account Field Char 10
Data 4 | User Account Field Char 11
Data 5 | User Account Field Char 12
Data 6 | User Account Field Char 13
Data 7 | User Account Field Char 14
Data 8 | User Account Field Char 15
Extended ID: | 0x10D4XYC3
Urgency: | 0x1 (Normal)
Function Code: | 0x0D4
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x3
DLC: | 0x8
Function Code: | 212
Frame: | 4 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Account Field Char 16
Data 2 | User Account Field Char 17
Data 3 | User Account Field Char 18
Data 4 | User Account Field Char 19
Data 5 | User Account Field Char 20
Data 6 | User Account Field Char 21
Data 7 | User Account Field Char 22
Data 8 | User Account Field Char 23
Extended ID: | 0x10D4XYC4
Urgency: | 0x1 (Normal)
Function Code: | 0x0D4
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x3 (Other Frame of Multi-Frame Message)
Frame Count: | 0x4
DLC: | 0x8
Function Code: | 212
Frame: | 5 of 5
| Upper Nibble |  |  |  | Lower Nibble
Data 1 | User Account Field Char 24
Data 2 | User Account Field Char 25
Data 3 | User Account Field Char 26
Data 4 | User Account Field Char 27
Data 5 | User Account Field Char 28
Data 6 | User Account Field Char 29
Data 7 | User Account Field Char 30
Data 8 | User Account Field Char 31
FC 400: Cloud Connection Status
Extended ID: | 0x1190XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x190
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x5
Function Code: | 400
Frame: | 1 of 1
Description: | Internet Connection Status - A message reportng the change to Internet Connection
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Connection to Cloud (False=0, True=1)
Data 2 | Connection to Cloud is Open (False=0, True=1)
Data 3 | SoftAP Status (Unknown=0, Unmanaged=10, Unavailable=20, Disconnected=30, Prepare=40, Config=50, NeedAuth=60, IpConfig=70, IpCheck=80, Secondaries=90, Activated=100, Deactivating=110, Failed=120)
Data 4 | Wireless Status (Unknown=0, Unmanaged=10, Unavailable=20, Disconnected=30, Prepare=40, Config=50, NeedAuth=60, IpConfig=70, IpCheck=80, Secondaries=90, Activated=100, Deactivating=110, Failed=120)
Data 5 | Wired Status (Unknown=0, Unmanaged=10, Unavailable=20, Disconnected=30, Prepare=40, Config=50, NeedAuth=60, IpConfig=70, IpCheck=80, Secondaries=90, Activated=100, Deactivating=110, Failed=120)
FC 403: Software Update Status
Extended ID: | 0x1193XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x193
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x2
Function Code: | 403
Description: | Software Update Status
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | #REF!
Data 2 | #REF!
FC 404: Software Update Scheduled Time
Extended ID: | 0x1194XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x194
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x4
Function Code: | 404
Frame: | 1 of 1
Description: | Software Update Scheduled Time - Scheduled Time for the software update, represents a uint32 as seconds since Epoch (default: 2016-01-01 00:00:00)
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Date/Time Stamp (Bits 0-7)
Data 2 | Date/Time Stamp (Bits 8-15)
Data 3 | Date/Time Stamp (Bits 16-23)
Data 4 | Date/Time Stamp (Bits 24-31)
FC 407: Device-Side Scheduled Firmware Update Interrupt
Extended ID: | 0x1197XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x197
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 1
Function Code: | 407
Frame: | 1 of 1
Description: | Sent to ack/nack the Software update status message when user action stage is reached. Also sent when user hits "update now" or "cancel"
Origin | Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | #REF! |  |  |  | #REF!
FC 408: Trigger Component Update
Extended ID: | 0x1198XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x198
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 1
Function Code: | 408
Frame: | 1 of 1
Description: | This message is sent from the product to the connectivity module to signal that the product component list in the device-twin should be updated.
Origin | Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Update Component (0 = Now)
FC 409: User Action Configuration Data
Extended ID: | 0x1199XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x199
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x7
Function Code: | 409
Frame: | 1 of 1
Description: | Sent as a response to FC404. Sending this message counts as an acknowledgement of the scheduled time and provides durations for certain parts of theupdate process
Origin | Connectivity Module
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Scheduled time Ack/Nack (0=ACK, 1=NACK)
Data 2 | In-use check time (minutes) (bits 0-7)
Data 3 | In-use check time (minutes) (bits 8-15)
Data 4 | Grace period (minutes) (bits 0-7)
Data 5 | Grace period (minutes) (bits 8-15)
Data 6 | User Action Duration (minutes) (bits 0-7)
Data 7 | User Action Duration (minutes) (bits 8-15)
FC 512: Data Frame
Extended ID: | 0x1200XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x200
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x3 - 0x8
Function Code: | 512
Frame: | 1 of 1
Description: | Data Frame - Multiple of this frame are sent in a sequence in order to send data larger than the max multi-frame message length provided by QMS10679. The first data byte is required, but subsequent data bytes are optional based on the overall data length. Refer to 081-10051-11 for usage details.
Origin | Connectivity Module, Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Sequence Number (bits 0-7)
Data 2 | Sequence Number (bits 8-15)
Data 3 | Data (Required)
Data 4 | Data (Optional)
Data 5 | Data (Optional)
Data 6 | Data (Optional)
Data 7 | Data (Optional)
Data 8 | Data (Optional)
FC 513: AckFrame
Extended ID: | 0x1201XY00
Urgency: | 0x1 (Normal)
Function Code: | 0x201
Node ID: | 0xXY (Varies upon sending node)
Frame Type: | 0x0 (Single Frame)
Frame Count: | 0x0
DLC: | 0x04
Function Code: | 513
Frame: | 1 of 1
Description: | Ack Frame - Used to Ack/Nack FC512 data frames. Refer to 081-10051-11 for usage.
Origin | Connectivity Module, Display
| Bit 7 | Bit 6 | Bit 5 | Bit 4 | Bit 3 | Bit 2 | Bit 1 | Bit 0
Data 1 | Sequence Base (bits 0-7)
Data 2 | Sequence Base (bits 8-15)
Data 3 | Node ID to ACK
Data 4 | Status (1=>ACK, 2=>NACK, 3=>CRC Success, 4=>CRC Fail)
Note on C197: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Should this be in tenths of degrees F, to stay consistent with function code 07?